ECG Pipeline — Validation, Problem-Animal Triage, and Open Questions

Status report for the supervisor meeting. Consolidates the checks run on the pipeline: confirmation it works on the clean animals, the filter/FFT verification, a triage of the 34 flagged animals, a recovery analysis, and point-by-point answers to the open questions. Every number is reproducible from the notebooks and the scripts in code/.

# Executive summary

- The pipeline works on the clean animals. 83 of 117 recordings are clean; heart rate, rhythm regularity, and QTc are all physiological. Beat detection is confirmed correct.

- The filter is verified. It removes the 50 Hz mains oscillation (Luke's “oscillation across the trace”) and out-of-band noise while preserving the heartbeat — shown in both the time domain and the FFT power spectrum.

- The 34 “problem” animals are a gradient, not a wall: ~6 are essentially clean, 1–2 are recoverable by relocating the analysis window, ~16 are irregular throughout the recording (a physiological question, not a data problem), and ~10 are genuinely too noisy to use.

- Two items are for the supervisors, not for data processing: whether the irregular-throughout animals reflect real arrhythmia (Roisin), and the scope question of re-examining every recording from scratch (Luke/Roisin).

# 1. The pipeline works on the clean animals (the headline)

The proper test of the pipeline is the clean animals, not the known problem cases.

| Metric | Result | Verdict |
| --- | --- | --- |
| Clean animals | 83 of 117 | the majority |
| Heart rate | 385–626 bpm (median 530) | all inside the 300–700 physiological gate |
| Rhythm (rr_cv) | median 0.025 | very regular — locking onto real beats |
| QTc | median 41.5 ms | matches mouse literature (~41 ms) |

Beat detection on a clean animal is shown in the validation overlay (overlay_validation_A201.png): every beat is detected, and the averaged-beat template shows clean P / QRS / J-wave / T morphology.

# 2. Filter verification (FFT power spectrum)

Two weeks ago Luke saw a visible oscillation running through the whole trace. That oscillation is 50 Hz mains hum. Re-running the FFT with the current settings confirms it is removed:

- 50 Hz mains: ~90% removed on clean animals, up to 98.6% on noisy ones.

- Everything above 150 Hz is cut by the bandpass.

- The evenly-spaced comb of peaks that remains is not noise — it is the heartbeat's harmonics (HR fundamental ~8.3 Hz = 499 bpm, plus overtones). A periodic heartbeat must produce that comb.

Overlaying raw vs filtered in the time domain confirms the same thing: on clean animals the two track each other closely (correlation 0.93–0.97) and only ~26–39% of the signal is removed (noise + drift), while the heartbeat shape is preserved.

# 3. The small bumps on clean animals are real, not noise

Between the R-peak and baseline on a clean beat there are small deflections. These are genuine cardiac features, not residual noise: the P wave, the J-wave (prominent in mice, immediately after R), and the inverted T-wave. Mice have no flat ST segment, so repolarisation begins almost immediately. These should be explained, not filtered away.

# 4. Measurement note — the QRS-duration field

The current qrs_duration_ms measures the FWHM of the R-spike (~6 ms), not the clinical QRS duration (Q-onset → S-offset, ~10–29 ms). It systematically under-reads the true QRS width. Because anthracycline cardiotoxicity widens the QRS, the definition matters — this is flagged for Roisin to decide the QRS-offset definition (genuinely ambiguous in mouse ECG, where the J-wave runs straight on with no clean isoelectric point). Interim: rename the field to r_width_fwhm_ms so it is not mistaken for QRS duration.

# 5. Triage of the 34 flagged animals

The 34 flagged recordings span a gradient, measured by raw-vs-filter correlation and the fraction of the signal that is noise:

| Group | Count | Criterion | Action |
| --- | --- | --- | --- |
| Barely-flagged (essentially clean) | 6 | corr 0.95–0.99, 17–32% removed | keep via a two-tier gate |
| Recoverable by window relocation | 1–2 | a cleaner window exists in the recording | move to clean set after review |
| Irregular throughout | ~16 | irregularity spread across the whole recording | physiological question — Roisin's call |
| Severe / mostly noise | ~10 | corr < 0.75 and/or > 80% is noise | drop |

Barely-flagged (6): 117, 129, 131, 134, 139, 245. Filter behaves perfectly; they only tripped the rr_cv > 0.15 gate on mild irregularity.

Severe (~10): 102, 103, 107, 109, 113, 115, 120, 125, 127, 240. Example — Animal 120: raw and filtered barely track (corr 0.39), 97% of the trace is noise, HR 93 bpm (impossible for an anaesthetised mouse), polarity inverted. Unusable.

# 6. Recovery analysis (what “recovery” can and cannot do)

Each LabChart file is ~29 minutes but only a short window is measured. If that window landed on a noisy stretch, relocating to the cleanest window in the same recording (quality_window_search) can recover usable beats. Result across the 18 middle-group animals:

- 1 clean recovery: Animal 123 — rr_cv 0.220 → 0.089, HR 503 → 502 bpm (consistent), now below the 0.15 gate.

- 1 borderline: Animal 126 — a clean window exists (rr_cv 0.585 → 0.047) but HR shifts 452 → 559 bpm, so it needs manual review.

- 16 of 18: irregular throughout. Relocating to the cleanest window does not fix the irregularity — meaning it is intrinsic to the animal's rhythm, not a windowing artifact.

This is the key finding: when relocation cannot fix the irregularity, the cause is either real physiology (sinus arrhythmia / tachycardia / bradycardia — a finding, and Roisin's call) or a genuinely broken recording. It cannot be recovered by data processing, and that is the correct, defensible conclusion.

How each recovery is checked against the original:

- Visual overlay — detected R-peaks sit on real raw beats (real beats?).

- Physiological sanity — HR 300–700 bpm, rr_cv < 0.15, QTc ~41 ms.

- HR stability — recovered HR matches the original HR (same animal, not a spurious signal). This is the decisive check.

- Manual cross-check — against LabChart cursor values where available.

# 7. Point-by-point answers to the open questions

| # | Question | Answer |
| --- | --- | --- |
| 1 | Do the FFT spectra look better with the new settings? | Yes — 50 Hz mains removed (~90%), out-of-band noise cut, cardiac harmonics preserved. |
| 2 | Problem animals 111/113 — placement? baseline? re-examine from scratch? | 113/120 flagged polarity-inverted (reversed electrodes plausible). Baseline drift handled by a bounded window. “Re-examine everything from scratch” is a scope change — a question for the supervisors, not a plan. |
| 3 | Are the small bumps on clean animals noise? | No — they are P / J-wave / T waves (real physiology). |
| 4 | Do the clean animals pass, with correct beat detection? | Yes — 83 animals, HR physiological, rr_cv 0.025, QTc 41.5 ms. |
| 5 | Did any treatment cause real tachy/bradycardia? | Roisin's physiological call — cannot be settled from traces alone, and needs the treatment labels. Some “fast clustered” traces may be real tachycardia rather than detection failure. |
| 6 | Clear, interpretable plots for clean and problem animals? | Delivered — overlay, FFT, and per-animal raw-vs-filtered figures. |

# 8. Open questions for the supervisors

- QRS-offset definition (Roisin) — the current field measures R-spike width, not clinical QRS duration; the correct definition is ambiguous in mouse ECG.

- Physiological call on irregular-throughout animals (Roisin) — real arrhythmia vs broken recording; needs the treatment metadata.

- Scope of re-examination (Luke/Roisin) — whether to re-check every recording from scratch rather than relying on the manual LabChart comments.

# Appendix — figures

- overlay_validation_A201.png, overlay_validation_A111.png — raw → filtered → features

- fft_overlay_A201.png — FFT power spectrum, before vs after (overlaid)

- clean_beat_qrs_A201.png — labelled clean beat + QRS FWHM vs true QRS

- clean_rawvsfilt_A133.png (and 122, 116) — clean-animal raw vs filtered

- problem_animals/ — raw vs filtered for all 34 flagged animals

- recovered/recovered_A123.png, recovered_A126.png — recovery checks